## DBS
[62 empty] | CRM Batch 171 | 2029.19 | 0.87 | 2217.4 | 0.63 | Salinity33.434
|  | FROM DBS FILE | [59 empty] | CRM Batch | CT | CT StDev | AT | AT StDev |  |  | 250uL HgCl2/250ml | 250uL HgCl2/250ml
|  | ENTER CORRECT VALUES | do not change anthing else in the yellow section! | [58 empty] | raw | -24.85 | 3.2824 | 0.92104 | 2.2696 |  |  | 0.1 | 0.1 | %
|  | CALCULATED | [59 empty] | adj | 0.33835 | 1.5852 | 0.55732 | 1.992 |  |  | 1.001 | 1.001
| Type | Bottle | Day | Month | Year | Depth | I.S.T. | Sal. | Counts | Runtime | CT | FactorCT | Blank | TCT | LAST CT | CMRCT | LAST AT | CRMAT | CRM# | AT | FactorAT | AT RMS | CalcID | Titrino | A/B | Pip.Vol. | Comments | Lat. | Long. | DOA | TOA | Cell ID | Year | Day | Month | True SAL | True TEMP | True CT pipvol | Working AT pipvol | True AT pipvol | Working AT acid strenght | True AT acid strength | Baseline | Is CRM | Is STD | Is Junk | Is SMP | CT GOOD? | AT GOOD? | DateTime | Working density during analysis | True density during analysis | Recalculated CT | Recalculated AT | CTFactor | ATFactor | smooth CTFactor | smooth ATFactor | CT CORRECTED | AT CORRECTED | for stats only | for stats only | counter #CRM | smooth CTFactor | smooth ATFactor | CT CORRECTED | AT CORRECTED | Comments | CT CORRECTED HgCl2 | AT CORRECTED HgCl2
1 | bottle | old SW1 | 0 | 0 | 0 | 0 | 4 | 35 | 209040 | 12 | 2238.23 | 1 | 50 | 0.5 | 0 | 2029.19 | 0 | 2217.4 | 154 | 0 |  | NaN | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-19 | 21:06 | C_May04-05_2005 | 2018 | 19 | 12 | 35 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 1 | 0 | 1 | 1 | 2019-07-12 21:06 | 1.0233 | 1.0233 | 2238.5686 | 0 | #N/A | #N/A | 0.98775 | 0.99978 | 2266.3227 | 0
2 | bottle | old SW2 | 0 | 0 | 0 | 0 | 4 | 35 | 208909 | 12 | 2236.83 | 1 | 50 | 0.5 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2314.89 |  | 9.6e-05 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-19 | 21:32 | C_May04-05_2005 | 2018 | 19 | 12 | 35 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 1 | 0 | 1 | 1 | 2019-07-12 21:32 | 1.0233 | 1.0233 | 2237.1617 | 2326.6889 | #N/A | #N/A | 0.98927 | 0.99978 | 2261.4349 | 2327.1956
3 | bottle | old SW3 | 0 | 0 | 0 | 0 | 4 | 35 | 208745 | 12 | 2235.07 | 1 | 50 | 1.1 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2312.35 |  | 0.000102 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-19 | 21:49 | C_May04-05_2005 | 2018 | 19 | 12 | 35 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 1 | 0 | 1 | 1 | 2019-07-12 21:49 | 1.0233 | 1.0233 | 2235.4004 | 2324.136 | #N/A | #N/A | 0.98927 | 0.99978 | 2259.6545 | 2324.6421
4 | bottle | old SW4 new acid | 0 | 0 | 0 | 0 | 4 | 35 | 208273 | 12 | 2230 | 1 | 50 | 1.6 | 0 | 2029.19 | 0 | 2217.4 | 154 | 0 |  | NaN | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-19 | 22:12 | C_May04-05_2005 | 2018 | 19 | 12 | 35 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 1 | 0 | 1 | 1 | 2019-07-12 22:12 | 1.0233 | 1.0233 | 2230.3313 | 0 | #N/A | #N/A | 0.98927 | 0.99978 | 2254.5304 | 0
5 | bottle | old SW5 new acid | 0 | 0 | 0 | 0 | 4 | 35 | 208170 | 12 | 2228.89 | 1 | 50 | 2.1 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2319.97 |  | 0.0001 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-19 | 22:33 | C_May04-05_2005 | 2018 | 19 | 12 | 35 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 1 | 0 | 1 | 1 | 2019-07-12 22:33 | 1.0233 | 1.0233 | 2229.2252 | 2331.7948 | #N/A | #N/A | 0.98927 | 0.99978 | 2253.4122 | 2332.3026
6 | bottle | old SW6 new acid | 0 | 0 | 0 | 0 | 4 | 35 | 207877 | 12 | 2225.75 | 1 | 50 | 2.6 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2317.42 |  | 0.000123 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-19 | 22:51 | C_May04-05_2005 | 2018 | 19 | 12 | 35 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 1 | 0 | 1 | 1 | 2019-07-12 22:51 | 1.0233 | 1.0233 | 2226.0785 | 2329.2318 | #N/A | #N/A | 0.98927 | 0.99978 | 2250.2313 | 2329.739
7 | bottle | lnsw new acid | 0 | 0 | 0 | 0 | 4 | 35 | 187167 | 12 | 2003.36 | 1 | 50 | 3.1 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2264.92 |  | 0.000131 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-19 | 23:07 | C_May04-05_2005 | 2018 | 19 | 12 | 35 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 1 | 0 | 1 | 1 | 2019-07-12 23:07 | 1.0233 | 1.0233 | 2003.6607 | 2276.4642 | #N/A | #N/A | 0.98927 | 0.99978 | 2025.4003 | 2276.9599
8 | bottle | Junk1_Wad | 0 | 0 | 0 | 0 | 4 | 28 | 207239 | 12 | 2230.41 | 1 | 50 | 0 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2304.58 |  | 4.7e-05 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-20 | 8:26 | C_Nov20-18_0811 | 2018 | 20 | 12 | 28 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 1 | 0 | 1 | 1 | 2019-08-12 08:26 | 1.0181 | 1.0181 | 2230.7367 | 2316.3264 | #N/A | #N/A | 0.98927 | 0.99978 | 2254.9401 | 2316.8308
9 | bottle | Junk2_Wad | 0 | 0 | 0 | 0 | 4 | 28 | 206598 | 12 | 2223.49 | 1 | 50 | 0.5 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2309.81 |  | 6.7e-05 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-20 | 8:43 | C_Nov20-18_0811 | 2018 | 20 | 12 | 28 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 1 | 0 | 1 | 1 | 2019-08-12 08:43 | 1.0181 | 1.0181 | 2223.8169 | 2321.583 | #N/A | #N/A | 0.98927 | 0.99978 | 2247.9452 | 2322.0886
10 | bottle | Junk3_Wad | 0 | 0 | 0 | 0 | 4 | 28 | 206644 | 12 | 2223.99 | 1 | 50 | 1 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2307.6 |  | 7.2e-05 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-20 | 8:59 | C_Nov20-18_0811 | 2018 | 20 | 12 | 28 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 1 | 0 | 1 | 1 | 2019-08-12 08:59 | 1.0181 | 1.0181 | 2224.3135 | 2319.3618 | #N/A | #N/A | 0.98927 | 0.99978 | 2248.4472 | 2319.8668
11 | bottle | Junk4_Wad | 0 | 0 | 0 | 0 | 4 | 28 | 206604 | 12 | 2223.56 | 1 | 50 | 1.5 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2309.04 |  | 7.9e-05 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-20 | 9:15 | C_Nov20-18_0811 | 2018 | 20 | 12 | 28 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 1 | 0 | 1 | 1 | 2019-08-12 09:15 | 1.0181 | 1.0181 | 2223.8817 | 2320.8091 | #N/A | #N/A | 0.98927 | 0.99978 | 2248.0107 | 2321.3145
12 | bottle | Nutslab1 | 1 | 0 | 0 | 0 | 4 | 35 | 188645 | 12 | 2019.23 | 1 | 50 | 2.1 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2263.45 |  | 0.000155 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-20 | 9:32 | C_Nov20-18_0811 | 2018 | 20 | 12 | 35 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-08-12 09:32 | 1.0233 | 1.0233 | 2019.5339 | 2274.9867 | #N/A | #N/A | 0.98927 | 0.99978 | 2041.4457 | 2275.4821 |  |  | -1 | 0.98968 | 1.0002 | 2040.5877 | 2274.422 | Different SUB | 2042.6282 | 2276.6965
13 | bottle | CRM#171-0399 | 666 | 0 | 0 | 0 | 4 | 33.434 | 187372 | 12 | 2007.88 | 1 | 50 | 2.5 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2206.84 |  | 0.000189 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-20 | 9:49 | C_Nov20-18_0811 | 2018 | 20 | 12 | 33.434 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 1 | 0 | 0 | 0 | 1 | 1 | 2019-08-12 09:49 | 1.0222 | 1.0222 | 2008.1838 | 2218.0882 | 0.98965 | 1.0003 | 0.98927 | 0.99978 | 2029.9725 | 2218.5712 | 2029.9725 | 2218.5712 | 0 | 0.98965 | 1.0003 | 2029.19 | 2217.4 |  | 2031.2192 | 2219.6174
14 | bottle | CRM#171-0399B | 666 | 0 | 0 | 0 | 4 | 33.434 | 187228 | 12 | 2006.34 | 1 | 50 | 3 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2204.51 |  | 0.000225 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-20 | 10:06 | C_Nov20-18_0811 | 2018 | 20 | 12 | 33.434 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 1 | 0 | 0 | 0 | 1 | 1 | 2019-08-12 10:06 | 1.0222 | 1.0222 | 2006.6355 | 2215.7463 | 0.98888 | 0.99925 | 0.98927 | 0.99978 | 2028.4075 | 2216.2288 | 2028.4075 | 2216.2288 | 1 | 0.98961 | 1.0004 | 2027.6962 | 2214.9214 |  | 2029.7239 | 2217.1363
15 | bottle | ROTTMPT50_2018008771 | 31 | 7 | 2018 | 0 | 4 | 33.42 | 196501 | 12 | 2106.05 | 1 | 50 | 3.5 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2351.55 |  | 0.000178 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-20 | 10:24 | C_Nov20-18_0811 | 2018 | 20 | 12 | 33.42 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-08-12 10:24 | 1.0221 | 1.0221 | 2106.3612 | 2363.5358 | #N/A | #N/A | 0.98927 | 0.99978 | 2129.2151 | 2364.0504 |  |  | 2 | 0.98958 | 1.0004 | 2128.5427 | 2362.5092 |  | 2130.6712 | 2364.8717
16 | bottle | ROTTMPT70_2018008778 | 12 | 9 | 2018 | 0 | 4 | 33.46 | 205352 | 12 | 2201.14 | 1 | 50 | 4 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2366.37 |  | 0.000111 | 1 |  | red (1) | 18.8599 | Sample Name corrected | 0 | 0 | 2018-12-20 | 10:41 | C_Nov20-18_0811 | 2018 | 20 | 12 | 33.46 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-08-12 10:41 | 1.0222 | 1.0222 | 2201.4636 | 2378.4313 | #N/A | #N/A | 0.98927 | 0.99978 | 2225.3494 | 2378.9492 |  |  | 3 | 0.98954 | 1.0005 | 2224.7241 | 2377.2507 |  | 2226.9488 | 2379.628
17 | bottle | TERSLG235_2018005747 | 14 | 8 | 2018 | 0 | 4 | 34.76 | 191846 | 12 | 2053.97 | 1 | 50 | 4.5 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2288.63 |  | 0.000244 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-20 | 10:57 | C_Nov20-18_0811 | 2018 | 20 | 12 | 34.76 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-08-12 10:57 | 1.0232 | 1.0232 | 2054.2755 | 2300.2951 | #N/A | #N/A | 0.98927 | 0.99978 | 2076.5643 | 2300.796 |  |  | 4 | 0.98951 | 1.0006 | 2076.0532 | 2299.0106 |  | 2078.1293 | 2301.3096
18 | bottle | WALCRN2_2018005580 | 12 | 9 | 2018 | 0 | 4 | 32 | 202861 | 12 | 2176.71 | 1 | 50 | 4.9 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2413.23 |  | 0.000171 | 1 |  | red (1) | 18.8599 | Estimated Salinity | 0 | 0 | 2018-12-20 | 11:20 | C_Nov20-18_0811 | 2018 | 20 | 12 | 32 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-08-12 11:20 | 1.0211 | 1.0211 | 2177.0278 | 2425.5302 | #N/A | #N/A | 0.98927 | 0.99978 | 2200.6485 | 2426.0583 |  |  | 5 | 0.98948 | 1.0006 | 2200.1835 | 2424.0254 |  | 2202.3837 | 2426.4494
19 | bottle | WALCRN20_2018005592 | 12 | 9 | 2018 | 0 | 4 | 35 | 200206 | 12 | 2143.37 | 1 | 50 | 5.5 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2398.98 |  | 0.000307 | 1 |  | red (1) | 18.8599 | Estimated Salinity | 0 | 0 | 2018-12-20 | 11:36 | C_Nov20-18_0811 | 2018 | 20 | 12 | 35 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-08-12 11:36 | 1.0233 | 1.0233 | 2143.6947 | 2411.2075 | #N/A | #N/A | 0.98927 | 0.99978 | 2166.9538 | 2411.7326 |  |  | 6 | 0.98944 | 1.0007 | 2166.5714 | 2409.5621 |  | 2168.738 | 2411.9717
20 | bottle | WALCRN70_2018005604 | 12 | 9 | 2018 | 0 | 4 | 32.5 | 194397 | 12 | 2084.85 | 1 | 50 | 6 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2315.3 |  | 0.000125 | 1 |  | red (1) | 18.8599 | Estimated Salinity | 0 | 0 | 2018-12-20 | 11:52 | C_Nov20-18_0811 | 2018 | 20 | 12 | 32.5 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-08-12 11:52 | 1.0215 | 1.0215 | 2085.1554 | 2327.101 | #N/A | #N/A | 0.98927 | 0.99978 | 2107.7792 | 2327.6077 |  |  | 7 | 0.98941 | 1.0007 | 2107.4808 | 2325.3687 |  | 2109.5883 | 2327.6941
21 | bottle | SCHOUWN10_2018005616 | 12 | 9 | 2018 | 0 | 4 | 31 | 200946 | 12 | 2157.69 | 1 | 50 | 6.4 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2416.93 |  | 0.000122 | 1 |  | red (1) | 18.8599 | Estimated Salinity | 0 | 0 | 2018-12-20 | 12:09 | C_Nov20-18_0811 | 2018 | 20 | 12 | 31 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-08-12 12:09 | 1.0203 | 1.0203 | 2158.01 | 2429.249 | #N/A | #N/A | 0.98927 | 0.99978 | 2181.4243 | 2429.778 |  |  | 8 | 0.98937 | 1.0008 | 2181.1914 | 2427.2901 |  | 2183.3726 | 2429.7174
22 | bottle | GOERE2_2018005628 | 12 | 9 | 2018 | 0 | 4 | 30.5 | 206478 | 12 | 2218.09 | 1 | 50 | 6.9 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2427.05 |  | 0.000124 | 1 |  | red (1) | 18.8599 | Estimated Salinity | 0 | 0 | 2018-12-20 | 12:26 | C_Nov20-18_0811 | 2018 | 20 | 12 | 30.5 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-08-12 12:26 | 1.0199 | 1.0199 | 2218.4171 | 2439.4206 | #N/A | #N/A | 0.98927 | 0.99978 | 2242.4868 | 2439.9518 |  |  | 9 | 0.98934 | 1.0009 | 2242.3257 | 2437.3023 |  | 2244.568 | 2439.7396
23 | bottle | GOERE6_2018005640 | 12 | 9 | 2018 | 0 | 4 | 30.5 | 206725 | 12 | 2220.75 | 1 | 50 | 7.5 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2429.62 |  | 0.000161 | 1 |  | red (1) | 18.8599 | Estimated Salinity | 0 | 0 | 2018-12-20 | 12:41 | C_Nov20-18_0811 | 2018 | 20 | 12 | 30.5 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-08-12 12:41 | 1.0199 | 1.0199 | 2221.0786 | 2442.0037 | #N/A | #N/A | 0.98927 | 0.99978 | 2245.1772 | 2442.5355 |  |  | 10 | 0.9893 | 1.0009 | 2245.0941 | 2439.7318 |  | 2247.3392 | 2442.1716
24 | bottle | NOORDWK2_2018005652 | 10 | 9 | 2018 | 0 | 4 | 30 | 204868 | 12 | 2201.56 | 1 | 50 | 8 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2414.47 |  | 0.000149 | 1 |  | red (1) | 18.8599 | Estimated Salinity | 0 | 0 | 2018-12-20 | 12:57 | C_Nov20-18_0811 | 2018 | 20 | 12 | 30 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-08-12 12:57 | 1.0196 | 1.0196 | 2201.8824 | 2426.7765 | #N/A | #N/A | 0.98927 | 0.99978 | 2225.7727 | 2427.3049 |  |  | 11 | 0.98927 | 1.001 | 2225.7679 | 2424.3684 |  | 2227.9937 | 2426.7928
25 | bottle | NOORDWK10_2018005664 | 10 | 9 | 2018 | 0 | 4 | 31 | 200341 | 12 | 2151.18 | 1 | 50 | 8.5 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2408.33 |  | 0.00015 | 1 |  | red (1) | 18.8599 | Estimated Salinity | 0 | 0 | 2018-12-20 | 13:14 | C_Nov20-18_0811 | 2018 | 20 | 12 | 31 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-08-12 13:14 | 1.0203 | 1.0203 | 2151.4932 | 2420.6052 | #N/A | #N/A | 0.98927 | 0.99978 | 2174.8369 | 2421.1323 |  |  | 12 | 0.98923 | 1.0011 | 2174.908 | 2418.0533 |  | 2177.0829 | 2420.4713
26 | bottle | NOORDWK20_2018005676 | 10 | 9 | 2018 | 0 | 4 | 31.5 | 198324 | 12 | 2128.67 | 1 | 50 | 9 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2393.61 |  | 0.000116 | 1 |  | red (1) | 18.8599 | Estimated Salinity | 0 | 0 | 2018-12-20 | 13:30 | C_Nov20-18_0811 | 2018 | 20 | 12 | 31.5 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-08-12 13:30 | 1.0207 | 1.0207 | 2128.9805 | 2405.8102 | #N/A | #N/A | 0.98927 | 0.99978 | 2152.0798 | 2406.334 |  |  | 13 | 0.9892 | 1.0011 | 2152.2252 | 2403.1248 |  | 2154.3775 | 2405.5279
27 | bottle | NOORDWK70_2018005688 | 10 | 9 | 2018 | 0 | 4 | 34.5 | 193804 | 12 | 2075.4 | 1 | 50 | 9.5 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2328.18 |  | 0.000245 | 1 |  | red (1) | 18.8599 | Estimated Salinity | 0 | 0 | 2018-12-20 | 13:46 | C_Nov20-18_0811 | 2018 | 20 | 12 | 34.5 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-08-12 13:46 | 1.023 | 1.023 | 2075.706 | 2340.0467 | #N/A | #N/A | 0.98927 | 0.99978 | 2098.2273 | 2340.5562 |  |  | 14 | 0.98917 | 1.0012 | 2098.4423 | 2337.2898 |  | 2100.5407 | 2339.6271
28 | bottle | TERSLG10_2018005700 | 11 | 9 | 2018 | 0 | 4 | 33.5 | 195484 | 12 | 2094.99 | 1 | 50 | 10 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2341.98 |  | 0.000214 | 1 |  | red (1) | 18.8599 | Estimated Salinity | 0 | 0 | 2018-12-20 | 14:03 | C_Nov20-18_0811 | 2018 | 20 | 12 | 33.5 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-08-12 14:03 | 1.0222 | 1.0222 | 2095.3024 | 2353.917 | #N/A | #N/A | 0.98927 | 0.99978 | 2118.0363 | 2354.4296 |  |  | 15 | 0.98913 | 1.0012 | 2118.3272 | 2350.998 |  | 2120.4455 | 2353.349
29 | bottle | TERSLG50_2018005712 | 11 | 9 | 2018 | 0 | 4 | 34 | 196034 | 12 | 2100.13 | 1 | 50 | 10.4 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2316.84 |  | 0.000178 | 1 |  | red (1) | 18.8599 | Estimated Salinity | 0 | 0 | 2018-12-20 | 14:19 | C_Nov20-18_0811 | 2018 | 20 | 12 | 34 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-08-12 14:19 | 1.0226 | 1.0226 | 2100.4398 | 2328.6489 | #N/A | #N/A | 0.98927 | 0.99978 | 2123.2295 | 2329.1559 |  |  | 16 | 0.9891 | 1.0013 | 2123.5951 | 2325.617 |  | 2125.7187 | 2327.9426
30 | bottle | TERSLG100_2018005724 | 11 | 9 | 2018 | 0 | 4 | 34.5 | 194879 | 12 | 2086.94 | 1 | 50 | 10.9 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2311.42 |  | 0.000173 | 1 |  | red (1) | 18.8599 | Estimated Salinity | 0 | 0 | 2018-12-20 | 14:35 | C_Nov20-18_0811 | 2018 | 20 | 12 | 34.5 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-08-12 14:35 | 1.023 | 1.023 | 2087.2554 | 2323.2012 | #N/A | #N/A | 0.98927 | 0.99978 | 2109.902 | 2323.7071 |  |  | 17 | 0.98906 | 1.0014 | 2110.3389 | 2320.0326 |  | 2112.4492 | 2322.3526
31 | bottle | TERSLG135_2018005736 | 11 | 9 | 2018 | 0 | 4 | 34.5 | 192405 | 12 | 2060.37 | 1 | 50 | 11.4 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2302.36 |  | 0.000227 | 1 |  | red (1) | 18.8599 | Estimated Salinity | 0 | 0 | 2018-12-20 | 14:51 | C_Nov20-18_0811 | 2018 | 20 | 12 | 34.5 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-08-12 14:51 | 1.023 | 1.023 | 2060.6757 | 2314.0951 | #N/A | #N/A | 0.98927 | 0.99978 | 2083.034 | 2314.599 |  |  | 18 | 0.98903 | 1.0014 | 2083.5379 | 2310.7956 |  | 2085.6215 | 2313.1064
32 | bottle | WALCRN2_2018005581 | 17 | 10 | 2018 | 0 | 4 | 32.55 | 203655 | 12 | 2184.36 | 1 | 50 | 11.9 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2409.97 |  | 0.000172 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-20 | 15:08 | C_Nov20-18_0811 | 2018 | 20 | 12 | 32.55 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-08-12 15:08 | 1.0215 | 1.0215 | 2184.686 | 2422.2535 | #N/A | #N/A | 0.98927 | 0.99978 | 2208.3897 | 2422.781 |  |  | 19 | 0.98899 | 1.0015 | 2209.001 | 2418.6499 |  | 2211.21 | 2421.0685
33 | bottle | WALCRN20_2018005593 | 17 | 10 | 2018 | 0 | 4 | 32.02 | 201876 | 12 | 2166.07 | 1 | 50 | 12.4 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2409.13 |  | 0.000168 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-20 | 15:24 | C_Nov20-18_0811 | 2018 | 20 | 12 | 32.02 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-08-12 15:24 | 1.0211 | 1.0211 | 2166.3938 | 2421.4093 | #N/A | #N/A | 0.98927 | 0.99978 | 2189.8991 | 2421.9365 |  |  | 20 | 0.98896 | 1.0016 | 2190.5817 | 2417.657 |  | 2192.7722 | 2420.0746
34 | bottle | WALCRN70_2018005605 | 17 | 10 | 2018 | 0 | 4 | 35.01 | 193670 | 12 | 2073.17 | 1 | 50 | 12.9 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2307.47 |  | 0.000237 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-20 | 15:40 | C_Nov20-18_0811 | 2018 | 20 | 12 | 35.01 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-08-12 15:40 | 1.0234 | 1.0234 | 2073.4852 | 2319.2311 | #N/A | #N/A | 0.98927 | 0.99978 | 2095.9824 | 2319.7361 |  |  | 21 | 0.98892 | 1.0016 | 2096.7089 | 2315.4936 |  | 2098.8056 | 2317.8091
35 | bottle | SCHOUWN10_2018005617 | 17 | 10 | 2018 | 0 | 4 | 33.25 | 200324 | 12 | 2147.42 | 1 | 50 | 13.4 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2404.14 |  | 0.000208 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-20 | 15:55 | C_Nov20-18_0811 | 2018 | 20 | 12 | 33.25 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-08-12 15:55 | 1.022 | 1.022 | 2147.7364 | 2416.3938 | #N/A | #N/A | 0.98927 | 0.99978 | 2171.0393 | 2416.92 |  |  | 22 | 0.98889 | 1.0017 | 2171.8675 | 2412.3503 |  | 2174.0394 | 2414.7626
36 | bottle | GOERE2_2018005629 | 17 | 10 | 2018 | 0 | 4 | 30.98 | 202987 | 12 | 2179.71 | 1 | 50 | 13.9 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2422.44 |  | 0.000179 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-20 | 16:12 | C_Nov20-18_0811 | 2018 | 20 | 12 | 30.98 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-08-12 16:12 | 1.0203 | 1.0203 | 2180.0266 | 2434.7871 | #N/A | #N/A | 0.98927 | 0.99978 | 2203.6798 | 2435.3173 |  |  | 23 | 0.98885 | 1.0017 | 2204.5974 | 2430.5621 |  | 2206.802 | 2432.9927
37 | bottle | GOERE6_2018005641 | 17 | 10 | 2018 | 0 | 4 | 31.31 | 202362 | 12 | 2172.44 | 1 | 50 | 14.4 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2418.85 |  | 0.000201 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-20 | 16:28 | C_Nov20-18_0811 | 2018 | 20 | 12 | 31.31 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-08-12 16:28 | 1.0206 | 1.0206 | 2172.7644 | 2431.1788 | #N/A | #N/A | 0.98927 | 0.99978 | 2196.3388 | 2431.7082 |  |  | 24 | 0.98882 | 1.0018 | 2197.33 | 2426.8097 |  | 2199.5273 | 2429.2365
38 | bottle | Nutslab2 | 1 | 0 | 0 | 0 | 4 | 35 | 190625 | 12 | 2040.49 | 1 | 50 | 14.9 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2254.78 |  | 8.4e-05 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-20 | 16:44 | C_Nov20-18_0811 | 2018 | 20 | 12 | 35 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-08-12 16:44 | 1.0233 | 1.0233 | 2040.7983 | 2266.2725 | #N/A | #N/A | 0.98775 | 1.0004 | 2066.1004 | 2265.3316 |  |  | 25 | 0.98879 | 1.0019 | 2063.9438 | 2262.0596 |  | 2066.0078 | 2264.3216
39 | bottle | CRM#171-0212 | 666 | 0 | 0 | 0 | 4 | 33.434 | 187201 | 12 | 2006.05 | 1 | 50 | 15.4 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2210.4 |  | 0.000199 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-20 | 16:59 | C_Nov20-18_0811 | 2018 | 20 | 12 | 33.434 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 1 | 0 | 0 | 0 | 1 | 1 | 2019-08-12 16:59 | 1.0222 | 1.0222 | 2006.3452 | 2221.6663 | 0.98874 | 1.0019 | 0.98775 | 1.0004 | 2031.2201 | 2220.7439 | 2031.2201 | 2220.7439 | 26 | 0.98875 | 1.0019 | 2029.1708 | 2217.3989 |  | 2031.1999 | 2219.6163
40 | bottle | JUNK5_WAD | 0 | 0 | 0 | 0 | 4 | 28 | 204533 | 12 | 2201.2 | 1 | 50 | 0 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2305.39 |  | 6.6e-05 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-21 | 8:31 | C_Nov21-18_0811 | 2018 | 21 | 12 | 28 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 1 | 0 | 1 | 1 | 2019-09-12 08:31 | 1.0181 | 1.0181 | 2201.5245 | 2317.1405 | #N/A | #N/A | 0.98775 | 1.0004 | 2228.8193 | 2316.1784
41 | bottle | JUNK6_WAD | 0 | 0 | 0 | 0 | 4 | 28 | 204897 | 12 | 2205.13 | 1 | 50 | 0.5 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2313.93 |  | 8.1e-05 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-21 | 8:48 | C_Nov21-18_0811 | 2018 | 21 | 12 | 28 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 1 | 0 | 1 | 1 | 2019-09-12 08:48 | 1.0181 | 1.0181 | 2205.454 | 2325.724 | #N/A | #N/A | 0.98775 | 1.0004 | 2232.7975 | 2324.7584
42 | bottle | JUNK7_WAD | 0 | 0 | 0 | 0 | 4 | 28 | 204396 | 12 | 2199.72 | 1 | 50 | 1 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2313.77 |  | 0.000117 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-21 | 9:04 | C_Nov21-18_0811 | 2018 | 21 | 12 | 28 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 1 | 0 | 1 | 1 | 2019-09-12 09:04 | 1.0181 | 1.0181 | 2200.0456 | 2325.5632 | #N/A | #N/A | 0.98775 | 1.0004 | 2227.322 | 2324.5976
43 | bottle | JUNK8_WAD | 0 | 0 | 0 | 0 | 4 | 28 | 204412 | 12 | 2199.9 | 1 | 50 | 1.5 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2314.34 |  | 0.000157 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-21 | 9:20 | C_Nov21-18_0811 | 2018 | 21 | 12 | 28 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 1 | 0 | 1 | 1 | 2019-09-12 09:20 | 1.0181 | 1.0181 | 2200.2183 | 2326.1361 | #N/A | #N/A | 0.98642 | 1.0003 | 2230.5181 | 2325.3582
44 | bottle | NUTSLAB3 | 1 | 0 | 0 | 0 | 4 | 35 | 190302 | 12 | 2037.02 | 1 | 50 | 2 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2254.62 |  | 7.9e-05 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-21 | 9:39 | C_Nov21-18_0811 | 2018 | 21 | 12 | 35 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-09-12 09:39 | 1.0233 | 1.0233 | 2037.3294 | 2266.1117 | #N/A | #N/A | 0.98642 | 1.0003 | 2065.3861 | 2265.3539 |  |  | -1 | 0.98768 | 0.99943 | 2062.7518 | 2267.4121 |  | 2064.8146 | 2269.6795
45 | bottle | CRM#171-0637 | 666 | 0 | 0 | 0 | 4 | 33.434 | 186971 | 12 | 2003.57 | 1 | 50 | 2.5 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2205.04 |  | 0.000176 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-21 | 9:55 | C_Nov21-18_0811 | 2018 | 21 | 12 | 33.434 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 1 | 0 | 0 | 0 | 1 | 1 | 2019-09-12 09:55 | 1.0222 | 1.0222 | 2003.8723 | 2216.279 | 0.98752 | 0.99949 | 0.98642 | 1.0003 | 2031.4681 | 2215.5378 | 2031.4681 | 2215.5378 | 0 | 0.98752 | 0.99949 | 2029.19 | 2217.4 |  | 2031.2192 | 2219.6174
46 | bottle | NOORDWK2_2018005653 | 15 | 10 | 2018 | 0 | 4 | 30.1 | 205202 | 12 | 2205 | 1 | 50 | 3 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2431.05 |  | 0.000171 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-21 | 10:12 | C_Nov21-18_0811 | 2018 | 21 | 12 | 30.1 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-09-12 10:12 | 1.0196 | 1.0196 | 2205.3197 | 2443.441 | #N/A | #N/A | 0.98642 | 1.0003 | 2235.6897 | 2442.6238 |  |  | 1 | 0.98737 | 0.99956 | 2233.527 | 2444.5107 |  | 2235.7605 | 2446.9552
47 | bottle | NOORDWK10_2018005665 | 15 | 10 | 2018 | 0 | 4 | 30.83 | 202627 | 12 | 2176.07 | 1 | 50 | 3.5 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2420.52 |  | 0.000182 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-21 | 10:27 | C_Nov21-18_0811 | 2018 | 21 | 12 | 30.83 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-09-12 10:27 | 1.0202 | 1.0202 | 2176.3901 | 2432.8573 | #N/A | #N/A | 0.98642 | 1.0003 | 2206.3618 | 2432.0437 |  |  | 2 | 0.98722 | 0.99963 | 2204.5674 | 2433.7569 |  | 2206.772 | 2436.1906
48 | bottle | NOORDWK20_2018005677 | 15 | 10 | 2018 | 0 | 4 | 31.34 | 200789 | 12 | 2155.46 | 1 | 50 | 4 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2417.96 |  | 0.000192 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-21 | 10:43 | C_Nov21-18_0811 | 2018 | 21 | 12 | 31.34 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-09-12 10:43 | 1.0206 | 1.0206 | 2155.7771 | 2430.2843 | #N/A | #N/A | 0.98642 | 1.0003 | 2185.4649 | 2429.4715 |  |  | 3 | 0.98707 | 0.9997 | 2184.0243 | 2431.0176 |  | 2186.2084 | 2433.4486
49 | bottle | NOORDWK70_2018005689 | 15 | 10 | 2018 | 0 | 4 | 34.1 | 195730 | 12 | 2096.7 | 1 | 50 | 4.5 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2333.5 |  | 0.000242 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-21 | 10:59 | C_Nov21-18_0811 | 2018 | 21 | 12 | 34.1 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-09-12 10:59 | 1.0227 | 1.0227 | 2097.0176 | 2345.3938 | #N/A | #N/A | 0.98642 | 1.0003 | 2125.8962 | 2344.6094 |  |  | 4 | 0.98691 | 0.99977 | 2124.8227 | 2345.942 |  | 2126.9476 | 2348.288
50 | bottle | TERSLG10_2018005701 | 16 | 10 | 2018 | 0 | 4 | 33.86 | 196246 | 12 | 2102.62 | 1 | 50 | 5 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2354.07 |  | 0.000214 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-21 | 11:15 | C_Nov21-18_0811 | 2018 | 21 | 12 | 33.86 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-09-12 11:15 | 1.0225 | 1.0225 | 2102.9357 | 2366.0686 | #N/A | #N/A | 0.98642 | 1.0003 | 2131.8958 | 2365.2774 |  |  | 5 | 0.98676 | 0.99983 | 2131.1481 | 2366.4609 |  | 2133.2793 | 2368.8273
51 | bottle | TERSLG50_2018005713 | 16 | 10 | 2018 | 0 | 4 | 34 | 193095 | 12 | 2068.54 | 1 | 50 | 5.5 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2292.61 |  | 0.000195 | 1 |  | red (1) | 18.8599 | Estimated Salinity | 0 | 0 | 2018-12-21 | 11:31 | C_Nov21-18_0811 | 2018 | 21 | 12 | 34 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-09-12 11:31 | 1.0226 | 1.0226 | 2068.8527 | 2304.2954 | #N/A | #N/A | 0.98642 | 1.0003 | 2097.3434 | 2303.5248 |  |  | 6 | 0.98661 | 0.9999 | 2096.9314 | 2304.5207 |  | 2099.0284 | 2306.8253
52 | bottle | TERSLG100_2018005725 | 16 | 10 | 2018 | 0 | 4 | 34.6 | 195373 | 12 | 2092.1 | 1 | 50 | 6 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2316.81 |  | 0.00019 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-21 | 11:48 | C_Nov21-18_0811 | 2018 | 21 | 12 | 34.6 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-09-12 11:48 | 1.023 | 1.023 | 2092.4082 | 2328.6187 | #N/A | #N/A | 0.98642 | 1.0003 | 2121.2233 | 2327.84 |  |  | 7 | 0.98646 | 0.99997 | 2121.1339 | 2328.6882 |  | 2123.2551 | 2331.0169
53 | bottle | TERSLG135_2018005725737 | 16 | 10 | 2018 | 0 | 4 | 34 | 195729 | 12 | 2096.85 | 1 | 50 | 6.4 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2310.68 |  | 0.000182 | 1 |  | red (1) | 18.8599 | Estimated Salinity | 0 | 0 | 2018-12-21 | 12:04 | C_Nov21-18_0811 | 2018 | 21 | 12 | 34 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-09-12 12:04 | 1.0226 | 1.0226 | 2097.1618 | 2322.4575 | #N/A | #N/A | 0.98642 | 1.0003 | 2126.0424 | 2321.6808 |  |  | 8 | 0.98631 | 1 | 2126.281 | 2322.3689 |  | 2128.4073 | 2324.6913
54 | bottle | NUTSLAB4 | 1 | 0 | 0 | 0 | 4 | 35 | 190234 | 12 | 2036.29 | 1 | 50 | 6.9 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2253.92 |  | 7.7e-05 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-21 | 12:20 | C_Nov21-18_0811 | 2018 | 21 | 12 | 35 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 0 | 1 | 1 | 1 | 2019-09-12 12:20 | 1.0233 | 1.0233 | 2036.5991 | 2265.4082 | #N/A | #N/A | 0.98642 | 1.0003 | 2064.6457 | 2264.6506 |  |  | 9 | 0.98615 | 1.0001 | 2065.1963 | 2265.1679 |  | 2067.2615 | 2267.433
55 | bottle | CRM#171-0644 | 666 | 0 | 0 | 0 | 4 | 33.434 | 186680 | 12 | 2000.45 | 1 | 50 | 7.4 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2208.95 |  | 0.000171 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-21 | 12:37 | C_Nov21-18_0811 | 2018 | 21 | 12 | 33.434 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 1 | 0 | 0 | 0 | 1 | 1 | 2019-09-12 12:37 | 1.0222 | 1.0222 | 2000.7434 | 2220.2089 | 0.98598 | 1.0013 | 0.98642 | 1.0003 | 2028.2962 | 2219.4665 | 2028.2962 | 2219.4665 | 10 | 0.986 | 1.0002 | 2029.1504 | 2219.8226 |  | 2031.1796 | 2222.0424
56 | bottle | CRM#171-0644B | 666 | 0 | 0 | 0 | 4 | 33.434 | 186635 | 12 | 1999.96 | 1 | 50 | 7.9 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2206.69 |  | 0.000178 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-21 | 12:57 | C_Nov21-18_0811 | 2018 | 21 | 12 | 33.434 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 1 | 0 | 0 | 0 | 1 | 1 | 2019-09-12 12:57 | 1.0222 | 1.0222 | 2000.2596 | 2217.9374 | 0.98574 | 1.0002 | 0.98642 | 1.0003 | 2027.8057 | 2217.1957 | 2027.8057 | 2217.1957 | 11 | 0.98585 | 1.0002 | 2028.973 | 2217.4008 |  | 2031.002 | 2219.6182
57 | bottle | WADD_JUNK | 0 | 0 | 0 | 0 | 4 | 28 | 199499 | 12 | 2146.87 | 1 | 50 | 8.3 | 0 | 2029.19 | 0 | 2217.4 | 154 | 2310.15 |  | 8.4e-05 | 1 |  | red (1) | 18.8599 |  | 0 | 0 | 2018-12-21 | 13:15 | C_Nov21-18_0811 | 2018 | 21 | 12 | 28 | 25 | 18.86 | 100 | 97.256 | 0.1 | 0.1023 | 50 | 0 | 0 | 1 | 0 | 1 | 1 | 2019-09-12 13:15 | 1.0181 | 1.0181 | 2147.1808 | 2321.9248 | #N/A | #N/A | 0.98642 | 1.0003 | 2176.7502 | 2321.1483 | [6 empty] | 2nd CRM used as not sure if TA pipette filled properly at the first crm?
Formulas:
DateTime = DATE([Year],[Day],[Month])+[TOA]

## CT & AT Results
ID | Date | Place |  | Salinity | CT CORRECTED | AT CORRECTED | CT Hg CORRECTED | AT Hg CORRECTED | Comments
| 2018-11-20 | Nuts Lab Sub | Nutslab1 | 35 | 2040.5877 | 2274.422 |  |  | Old SUB. Refilled after this measurement.
| 2018-11-20 | CRM#171 | CRM#171-0399 | 33.434 | 2029.19 | 2217.4
2018008771 | 2018-07-31 07:14 | ROTTMPT50 | ROTTMPT50_2018008771 | 33.42 | 2128.5427 | 2362.5092 | 2130.6712 | 2364.8717 |  |  | NEW CT
2018008778 | 2018-07-31 08:46 | ROTTMPT70 | ROTTMPT70_2018008778 | 33.46 | 2224.7241 | 2377.2507 | 2226.9488 | 2379.628 | DIC High? Need to Recalcualte with CO2Sys | CT Recalcuated with CO2Sys (AT & pH) | 2147.29
2018005747 | 2018-08-14 13:39 | TERSLG235 | TERSLG235_2018005747 | 34.76 | 2076.0532 | 2299.0106 | 2078.1293 | 2301.3096
2018005580 | 2018-09-12 11:06 | WALCRN2 | WALCRN2_2018005580 | 32.5 | 2200.1835 | 2424.0254 | 2202.3837 | 2426.4494 | Salinity measured with NIOZ Conductivity Probe
2018005592 | 2018-09-12 09:45 | WALCRN20 | WALCRN20_2018005592 | 33.3 | 2166.5714 | 2409.5621 | 2168.738 | 2411.9717 | Salinity measured with NIOZ Conductivity Probe
2018005604 | 2018-09-12 06:37 | WALCRN70 | WALCRN70_2018005604 | 35.1 | 2107.4808 | 2325.3687 | 2109.5883 | 2327.6941 | Salinity measured with NIOZ Conductivity Probe
2018005616 | 2018-09-12 13:14 | SCHOUWN10 | SCHOUWN10_2018005616 | 32.8 | 2181.1914 | 2427.2901 | 2183.3726 | 2429.7174 | Salinity measured with NIOZ Conductivity Probe
2018005628 | 2018-09-12 15:11 | GOERE2 | GOERE2_2018005628 | 31.1 | 2242.3257 | 2437.3023 | 2244.568 | 2439.7396 | Salinity measured with NIOZ Conductivity Probe
2018005640 | 2018-09-12 15:39 | GOERE6 | GOERE6_2018005640 | 31 | 2245.0941 | 2439.7318 | 2247.3392 | 2442.1716 | Salinity measured with NIOZ Conductivity Probe
2018005652 | 2018-09-10 14:00 | NOORDWK2 | NOORDWK2_2018005652 | 30.2 | 2225.7679 | 2424.3684 | 2227.9937 | 2426.7928 | Salinity measured with NIOZ Conductivity Probe
2018005664 | 2018-09-10 14:40 | NOORDWK10 | NOORDWK10_2018005664 | 31.7 | 2174.908 | 2418.0533 | 2177.0829 | 2420.4713 | Salinity measured with NIOZ Conductivity Probe
2018005676 | 2018-09-10 15:37 | NOORDWK20 | NOORDWK20_2018005676 | 33 | 2152.2252 | 2403.1248 | 2154.3775 | 2405.5279 | Salinity measured with NIOZ Conductivity Probe
2018005688 | 2018-09-10 18:35 | NOORDWK70 | NOORDWK70_2018005688 | 34.6 | 2098.4423 | 2337.2898 | 2100.5407 | 2339.6271 | Salinity measured with NIOZ Conductivity Probe
2018005700 | 2018-09-11 15:14 | TERSLG10 | TERSLG10_2018005700 | 33.8 | 2118.3272 | 2350.998 | 2120.4455 | 2353.349 | Salinity measured with NIOZ Conductivity Probe
2018005712 | 2018-09-11 12:29 | TERSLG50 | TERSLG50_2018005712 | 34.4 | 2123.5951 | 2325.617 | 2125.7187 | 2327.9426 | Salinity measured with NIOZ Conductivity Probe
2018005724 | 2018-09-11 08:40 | TERSLG100 | TERSLG100_2018005724 | 34.6 | 2110.3389 | 2320.0326 | 2112.4492 | 2322.3526 | Salinity measured with NIOZ Conductivity Probe
2018005736 | 2018-09-11 06:05 | TERSLG135 | TERSLG135_2018005736 | 34.7 | 2083.5379 | 2310.7956 | 2085.6215 | 2313.1064 | Salinity measured with NIOZ Conductivity Probe
2018005581 | 2018-10-17 10:59 | WALCRN2 | WALCRN2_2018005581 | 32.55 | 2209.001 | 2418.6499 | 2211.21 | 2421.0685
2018005593 | 2018-10-17 09:45 | WALCRN20 | WALCRN20_2018005593 | 32.02 | 2190.5817 | 2417.657 | 2192.7722 | 2420.0746
2018005605 | 2018-10-17 06:45 | WALCRN70 | WALCRN70_2018005605 | 35.01 | 2096.7089 | 2315.4936 | 2098.8056 | 2317.8091
2018005617 | 2018-10-17 12:42 | SCHOUWN10 | SCHOUWN10_2018005617 | 33.25 | 2171.8675 | 2412.3503 | 2174.0394 | 2414.7626
2018005629 | 2018-10-17 14:20 | GOERE2 | GOERE2_2018005629 | 30.98 | 2204.5974 | 2430.5621 | 2206.802 | 2432.9927
2018005641 | 2018-10-17 14:54 | GOERE6 | GOERE6_2018005641 | 31.31 | 2197.33 | 2426.8097 | 2199.5273 | 2429.2365
| 2018-11-20 | Nuts Lab Sub | Nutslab2 | 35 | 2063.9438 | 2262.0596
| 2018-11-20 | CRM#171 | CRM#171-0212 | 33.434 | 2029.1708 | 2217.3989
| 2018-11-21 | Nuts Lab Sub | NUTSLAB3 | 35 | 2062.7518 | 2267.4121
| 2018-11-21 | CRM#171 | CRM#171-0637 | 33.434 | 2029.19 | 2217.4
2018005653 | 2018-10-15 13:21 | NOORDWK2 | NOORDWK2_2018005653 | 30.1 | 2233.527 | 2444.5107 | 2235.7605 | 2446.9552
2018005665 | 2018-10-15 14:54 | NOORDWK10 | NOORDWK10_2018005665 | 30.83 | 2204.5674 | 2433.7569 | 2206.772 | 2436.1906
2018005677 | 2018-10-15 15:40 | NOORDWK20 | NOORDWK20_2018005677 | 31.34 | 2184.0243 | 2431.0176 | 2186.2084 | 2433.4486
2018005689 | 2018-10-15 19:05 | NOORDWK70 | NOORDWK70_2018005689 | 34.1 | 2124.8227 | 2345.942 | 2126.9476 | 2348.288
2018005701 | 2018-10-16 16:14 | TERSLG10 | TERSLG10_2018005701 | 33.86 | 2131.1481 | 2366.4609 | 2133.2793 | 2368.8273
2018005713 | 2018-10-16 13:49 | TERSLG50 | TERSLG50_2018005713 | 34.3 | 2096.9314 | 2304.5207 | 2099.0284 | 2306.8253 | Salinity measured with NIOZ Conductivity Probe
2018005725 | 2018-10-16 10:24 | TERSLG100 | TERSLG100_2018005725 | 34.6 | 2121.1339 | 2328.6882 | 2123.2551 | 2331.0169
2018005737 | 2018-10-16 08:09 | TERSLG135 | TERSLG135_2018005725737 | 34.7 | 2126.281 | 2322.3689 | 2128.4073 | 2324.6913 | Salinity measured with NIOZ Conductivity Probe
| 2018-11-21 | Nuts Lab Sub | NUTSLAB4 | 35 | 2065.1963 | 2265.1679
| 2018-11-21 | CRM#171 | CRM#171-0644 | 33.434 | 2029.1504 | 2219.8226 |  |  | Unsure if TAlk Burrette filled properly?
| 2018-11-21 | CRM#171 | CRM#171-0644 | 33.434 | 2028.973 | 2217.4008 |  |  | TAlk of CRM#171-B used.
Precision (Nov 2018) VINDTA
| INTERNAL LAB REFERENCE OCEAN WATER 300L |  |  | Salinity | CT CORRECTED | AT CORRECTED
| 2018-11-20 | Nuts Lab Sub | Nutslab2 | 35 | 2063.9438 | 2262.0596
| 2018-11-21 | Nuts Lab Sub | NUTSLAB3 | 35 | 2062.7518 | 2267.4121
| 2018-11-21 | Nuts Lab Sub | NUTSLAB4 | 35 | 2065.1963 | 2265.1679
|  |  |  | Average | 2063.964 | 2264.8798
|  |  |  | Stdev | 1.2224 | 2.6878
|  |  |  | Precision | 1.2µM/kg | 2.7µM/kg |  |  | offset CT and AT set at 0 bias to Dickson standard seawater (CRM#171)